## Slide 1: EPI-Lo Sensor
Don Mitchell, Matt Hill, Ralph McNutt
DRAFT 2013-10-14 5:53 PM
Notes: introduction

## Slide 2: Outline EPI-Lo Sensor Design
ISIS Science Requirements Relevant for EPI-Lo
EPI-Lo Instrument
Block Diagram(s)
Principles of Operation Overview
Sensor Voltages
Electrostatic Optics
Collimators and Start Foils
Solid State Detectors & Stop Foils
  Anti-coincidence System
  GEANT
Microchannel Plate (MCP)
Mass Resolution
Light & Dust Mitigation
Follow-up from peer reviews
Summary
THIS IS NOT UP TO DATE...IGNORE THIS PAGE.

## Slide 3: The EPI-Lo Instrument Requirements
Parameter | Required | Goal (Capability) | Comment/Heritage
Electron Energies | 50 - 500keV | 25 - 1000 keV | Electron capability from JEDI, RBSPICE
Ion Energies | 50 keV/nucleon – 15000 keV Total E | 50 keV/nucleon – 15000 keV Total E | Capability partially based on RBSPICE capabilities. Top energy ~250keV/nuc for Fe
Energy Resolution | 45% for required energy range | 40% for required energy range | Telemetry limited
Time sampling | 5 sec | 1 sec | Telemetry and/or statistics limited
Angle resolution | <30° x <30° | Ions, ~15 x 12 to <30° x <30° e-, 45° | Varies with elevation
Pitch Angle (PA) Coverage | 0°-90° or 90°-180°, some samples in both hemispheres | 0°-90° or 90°-180°, some samples in both hemispheres | 
Time for Full PA | 1 – 5 sec | 1 – 5 sec | Telemetry limited
Ion Composition | H, He3, He4, C, O, Ne, Mg, Si, Fe | H, He3, He4, C, O, Ne, Mg, Si, Fe | He3/He4 ~50 to 1000 keV/nuc
Electron Sensitivity: j=Intensity | j = 1E1-1E6/cm2-s-sr | Sensor-G:0.144 (cm2.sr) Pixel-G: ~0.02 (cm2.sr) Up to 6E6 1/s counting | j=Intensity (1/cm2-s-sr) G=Geom. Factor (cm2-sr) 8 pixels/sensor
Ion Sensitivity | j = 1E1-1E6/cm2-s-sr | Sensor-G:0.16 (cm2.sr) Pixel-G: ~0.002 (cm2.sr) Up to 3.5E6/s rate (TOFxE) | 80 pixels/sensor

## Slide 4: The EPI-Lo Instrument

## Slide 5: EPI-Lo Instrument, cut-away
EPI-LO CROSS-SECTION VIEW ISOMETRIC
POWER SUPPLY ASSY
EVENT PWB ASSY
ANODE PWB ASSY

## Slide 6: The EPI-Lo Field of View
SUN
RAM

## Slide 7: Collimator Foil FOV
Foil FOVs shown in yellow (these are not with the latest collimators).

## Slide 8: The EPI-Lo Field of View
SUN
These trapezoids, with the rounded corners, represent EPI-Lo unobstructed FOV’s
This line(s) represents the obstruction caused by the TPS
This line(s) represent the obstruction caused by WISPR: Red was March 2013 and 2011-ish envelope and yellow is the Sept 2013 one. Outer purple is the closed door (Sept 2013).
Don’t think this red obstruction was in 2011 model. Part of WISPR?
At 0 degrees elevation, the viewing angle is 9.4 degrees off from looking straight at the sun
ion
electron

## Slide 9: EPI-Lo Wedge Components
EPI-LO INDIVIDUAL SENSOR WEDGE ASSY CROSS SECTION
BOTTOM MCP 2.9kV
TOP MCP, 900V
MCP GRID, 1kV
SCREWS, TORLON
SENSOR WEDGE COVER, GROUND
APERTURE START FOIL GRID, GROUND
INSULATOR, SENSOR WEDGE COVER
ULTEM 1000
MCP COMPRESSION PLATE, 1kV
(TOP OF ANODE PWB NOT SHOWN, 3KV)
COLLIMATORS, AT GROUND
ACCELERATION SURFACE, +1kV

## Slide 10: EPI-Lo Wedge, Exploded View
EPI-LO INDIVIDUAL SENSOR WEDGE ASSY CROSS SECTION
APERTURE START FOILS, GROUND
WEDGE COVER INSULATOR, ULTEM 2300
WEDGE GRID, 1KV
WEDGE COVER, GROUND
SIDE WALL REMOVED FOR CLARITY

## Slide 11: Sensor Functional Schematic
Microchannel
Plate (MCP)
Anode
Solid State
Detectors (SSDs)
Start Foils
Stop Foil
Start Foils
Collimators
Baffles
MCP Grid
Start Acceleration
Grid
Stop Accel. Grid
MCP Spacer

## Slide 12: Sensor Voltages
Ground
1000 V
Equipotential
Volume
900 V
2900 V
3000 V
start e-
acceleration
region
MCP
Anode
Ground
stop e-
accel.
region
SSDs
All Foils
at Ground
100 V retarding
potential region

## Slide 13: Secondary Electrons & Potentials
100 V retarding
potential region
Equipotential
Volume
2000 V
Across Plates
100 V
MCP-to-Anode
Grid
Grid
Grid
SSDs
start e-
acceleration
region
stop e-
accel.
region
MCP
Anode

## Slide 14: Ions (Energy, TOF, and Position)
Incident Ion
Secondary
Electrons
Start
Stop
MCP
Anode
SSDs

## Slide 15: Electrons: Energy (& Position)
Secondary
Electrons
(infrequent)
MCP
Anode
SSDs
Incident Electron

## Slide 16: Energetic Electron Measurement
Secondary Electrons from Start Foils possible, but Low Probability;
Start Electron present identifies entrance aperture
?
?
?
?
?

## Slide 17: Energetic Ion Measurement
Ion Measurement Logic can require TOF or not, but with no TOF, no species

## Slide 18: Electrostatic Optics
SIMION simulation of electron optics using ~3eV random initial electron velocities.

## Slide 19: Sensor Voltages
Ground
1000 V
Equipotential
Volume
900 V
2900 V
3000 V
start e-
acceleration
region
MCP
Anode
Ground
stop e-
accel.
region
SSDs
All Foils
at Ground
100 V retarding
potential region

## Slide 20
Elevation 1
Elevation 2
Elevation 3
Elevation 4
Elevation 5
Name | Mean TOF (ns) | TOF Sigma (ns)
Elevation 1 | 1.06 | 0.19
Elevation 2 | 2.17 | 0.13
Elevation 3 | 2.35 | 0.05
Elevation 4 | 2.33 | 0.06
Elevation 5 | 2.32 | 0.12
Stops | 5.23 | 0.11
Initial Energy Distribution
Electron Dispersion
Hit Locations
Sample Trajectories
SIMION, Allegrini Energy Distribution
20

## Slide 21: EPI-Lo Wedge, Exploded View
EPI-LO INDIVIDUAL SENSOR WEDGE ASSY CROSS SECTION
APERTURE START FOILS, GROUND
WEDGE COVER INSULATOR, ULTEM 2300
WEDGE GRID, 1KV
WEDGE COVER, GROUND
SIDE WALL REMOVED FOR CLARITY

## Slide 22: MCP Inter-Plate Spacer also a mask
EXPLODED VIEW “NEW” MCP ASSEMBLY
2.9KV LUG ASSY
900 V LUG ASSY
1KV LUG ASSY
GUARD WASHER
SCREW
WASHER
LINEAR EXPANDERS, 3 PLACES
2.9KV ELECTRODE
SLIDES FOR EXPANDERS, 2 PLACES
SLIDE-SPACER-ELECTRODE
MCP
SPACER .002
ALTERNATE SPACER
MCP
900V ELECTRODE
COMPRESSION PLATE
INSULATED SHIM, ULTEM 1000
NUT
SCREW

## Slide 23
Elevation 1
Elevation 2
Elevation 3
Elevation 4
Elevation 5
Name | Mean TOF (ns) | TOF Sigma (ns)
Elevation 1 | 1.06 | 0.19
Elevation 2 | 2.17 | 0.13
Elevation 3 | 2.35 | 0.05
Elevation 4 | 2.33 | 0.06
Elevation 5 | 2.32 | 0.12
Stops | 5.23 | 0.11
Initial Energy Distribution
Electron Dispersion
Hit Locations
Sample Trajectories
SIMION, Alegrini Energy Distribution
23

## Slide 24: Microchannel Plate (MCP)

## Slide 25: Mass Resolution

## Slide 26: Mass Resolution

## Slide 27: Collimators and Start Foils
Foils Mount on Apertures (red)
    Each elevation tailored for equal GF
    Second foil planned at intermediate baffle (vented)
    Third foil may be needed for SW electrons
Collimators screw-mount to outer cover, capture foils
    Each Elevation and azimuth is unique

## Slide 28: EPI-Lo SSD Assemblies (8)
SSD ASSY TO WEDGE EXPLODED CROSS SECTION
ENERGY PWB ASSY
SSD ASSY
SSD ASSY HOUSING, ULTEM 1000
SSD WEDGE GRID, 1KV
SSD STOP GRID-FOIL, GROUND
SSD INSULATOR, ULTEM 1000
SCREWS, TORLON
GROUNDING SCREW
SENSOR WEDGE 1KV

## Slide 29: Anti-coincidence System
Electron SSD is backed by an anti-coincidence SSD

## Slide 30: Anti-coincidence System: GEANT
Setup for GEANT4 model of electron SSD anticoincidence

## Slide 31: EPI-Lo Block Diagram

## Slide 32: Sensor Voltages
MCP
Anode/Preamp
MCP
+2900V
HVPS Out +3000V
+ 900V
MCP Return is a ground reference at HVPS (allows MCP current measurement across resistor)
SSD
900V (total diode drop)
100
+3000V
1000V
Grounded Foils
Collimators
MCP return
Tecan Grid
100V
Voltage Control Cartoon

## Slide 33: EPI-Lo Block Diagram

## Slide 34: Light & Dust Mitigation
Dust may produce pinholes in the Start and collimator foils. Foils are designed to reduce UV by ~3 orders of magnitude. Pinholes may account for as much as ~3.5% of a foil area, so one pinhole can reduce effectiveness of attenuation to ~1/30.
For the 4 foils closest to the TPS edge, the suppression factor must be ~4 orders of magnitude. For these, pinholes are even more problematic.
Intensity of Ly-a EUV vs elongation. Most EPI-Lo entrances include angles no closer than 20 degrees elongation (which implies a maximum Ly-a intensity of ~109). For 4 apertures close to the TPS, the maximum is ~1012, but average ~1010.
Edge of TPS

## Slide 35: Light & Dust Mitigation
Dust may produce pinholes in the Start and collimator foils. Foils are designed to stop solar wind energy electrons. Pinholes may account for as much as ~3.5% of a foil area, so one pinhole can reduce effectiveness of attenuation to ~1/30.
Solar wind electrons have an energy of ~100eV or less. If these enter the Start foils (through a pin hole), the electron optics will nominally guide them to the Start areas of the MCP.
The solar wind electron flux is ~2 x 1012/cm2-s-sr. A foil pinhole, if it includes the full area of a support grid element, is ~6.25 x 10-4 cm2. The angular acceptance of a collimator is ~.02 sr. So the geometric factor of a pinhole is ~1.25 x 10-5cm2sr. So the flux through a pinhole would be expected to be ~2.5 x 107/s. If every aperture had 1 pinhole, the total for a quadrant would be ~5 x 108/s, clearly too much.
Adding a second foil in each aperture would greatly reduce the number of solar wind electrons expected to enter the TOF volume, but would still result in ~5 x 106/s, a bit more than we would like.
Mitigation might involve either a third foil, or a design for which pinholes are either much smaller than a grid element, or low probability.

## Slide 36: Light & Dust Mitigation
Calculation of impact of pinholes on rates. Because photoelectron efficiency is <10%, the UV rates are not a problem. The solar wind electron rates could be a problem.

## Slide 37: Light & Dust Mitigation
TBD

## Slide 38: Dust: Simulating the Environment
+Z-ish
45 deg
90 deg
135 deg
180 deg
225 deg
270 deg
315 deg
azimuthal angle
0 deg
0 deg
22.5 deg
45 deg
67.5 deg
90 deg
+Z
+x-ish
elevation angle
align with plane of instrument
Virtual surface
Back foil
Collimator
opening
Impacts within the back foil angle will pass through the collimator opening and front foil and hit the back foil
Front foil
Impacts within the front foil angle will pass through the collimator opening and hit the front foil
Impacts in the hemisphere will hit the “virtual surfaces” (ignores the collimator)
NOTE: actually only polar angle is limited, not position on surface, so foil area was adjusted to account for this and produce actual foil geometry factors,

## Slide 39: Dust: Simulating the Environment
Each unprotected foil would expect ~10 damage-inducing hits per mission. Reduced to ~1 hit per mission with collimator. Order of magnitude decrease.

## Slide 40: Dust: Lab tests
TBD

## Slide 41: SLIDE PREPARATION NOTES
Are we handling dust in this presentations? Assuming yes.
Should we discuss photons suppression here? Assuming yes, it’s background.
How about details like how the MCPs work (in general)? Or this could be handled when someone discusses the anode boards. Special features like the inter-plate mask (or is that handled in a mechanical section)?
Backup slides: MCPs
We are only supposed to respond to requirements, so I wasn’t going to go into any of the rate estimates from the work that I (Matt) and Rob D. did since that is for goals. Is this the right approach?

## Slide 42: Follow Up from Peer Reviews
TBD

## Slide 43: Summary
TBD

## Slide 44: Backup

## Slide 45: EPI-Lo Instrument Peer Review
EPI-LO INDIVIDUAL SENSOR WEDGE ASSY CROSS SECTION
APERTURE START FOILS, GROUND
WEDGE COVER INSULATOR, ULTEM 2300
WEDGE GRID, 1KV
WEDGE COVER, GROUND
SIDE WALL REMOVED FOR CLARITY

## Slide 46: EPI-Lo Instrument Peer Review
EXPLODED VIEW “NEW” MCP ASSEMBLY
2.9KV LUG ASSY
900 V LUG ASSY
1KV LUG ASSY
GUARD WASHER
SCREW
WASHER
LINEAR EXPANDERS, 3 PLACES
2.9KV ELECTRODE
SLIDES FOR EXPANDERS, 2 PLACES
SLIDE-SPACER-ELECTRODE
MCP
SPACER .002
ALTERNATE SPACER
MCP
900V ELECTRODE
COMPRESSION PLATE
INSULATED SHIM, ULTEM 1000
NUT
SCREW

## Slide 47: Principles of Operation