## Slide 1: EPI-Lo Sensor
Don Mitchell, Matt Hill, Ralph McNutt
DRAFT 2013-10-14 5:53 PM
Notes: introduction

## Slide 2: Outline EPI-Lo Sensor Design
ISIS Science Requirements Relevant for EPI-Lo
EPI-Lo Instrument
Block Diagram(s)
Principles of Operation Overview
Sensor Voltages
Electrostatic Optics
Collimators and Start Foils
Solid State Detectors & Stop Foils
  Anti-coincidence System
  GEANT
Microchannel Plate (MCP)
Mass Resolution
Light & Dust Mitigation
Follow-up from peer reviews
Summary
THIS IS NOT UP TO DATE...IGNORE THIS PAGE.

## Slide 3: The EPI-Lo Instrument Requirements
Parameter | Required | Goal (Capability) | Comment/Heritage
Electron Energies | 50 - 500keV | 25 - 1000 keV | Electron capability from JEDI, RBSPICE
Ion Energies | 50 keV/nucleon – 15000 keV Total E | 50 keV/nucleon – 15000 keV Total E | Capability partially based on RBSPICE capabilities. Top energy ~250keV/nuc for Fe
Energy Resolution | 45% for required energy range | 40% for required energy range | Telemetry limited
Time sampling | 5 sec | 1 sec | Telemetry and/or statistics limited
Angle resolution | <30° x <30° | Ions, ~15 x 12 to <30° x <30° e-, 45° | Varies with elevation
Pitch Angle (PA) Coverage | 0°-90° or 90°-180°, some samples in both hemispheres | 0°-90° or 90°-180°, some samples in both hemispheres | 
Time for Full PA | 1 – 5 sec | 1 – 5 sec | Telemetry limited
Ion Composition | H, He3, He4, C, O, Ne, Mg, Si, Fe | H, He3, He4, C, O, Ne, Mg, Si, Fe | He3/He4 ~50 to 1000 keV/nuc
Electron Sensitivity: j=Intensity | j = 1E1-1E6/cm2-s-sr | Sensor-G:0.144 (cm2.sr) Pixel-G: ~0.02 (cm2.sr) Up to 6E6 1/s counting | j=Intensity (1/cm2-s-sr) G=Geom. Factor (cm2-sr) 8 pixels/sensor
Ion Sensitivity | j = 1E1-1E6/cm2-s-sr | Sensor-G:0.16 (cm2.sr) Pixel-G: ~0.002 (cm2.sr) Up to 3.5E6/s rate (TOFxE) | 80 pixels/sensor

## Slide 4: The EPI-Lo Instrument

## Slide 5: EPI-Lo Instrument, cut-away
EPI-LO CROSS-SECTION VIEW ISOMETRIC
POWER SUPPLY ASSY
EVENT PWB ASSY
ANODE PWB ASSY

## Slide 6: The EPI-Lo Field of View
SUN
RAM

## Slide 7: The EPI-Lo Field of View
SUN
These trapezoids, with the rounded corners, represent EPI-Lo unobstructed FOV’s
This line(s) represents the obstruction caused by the TPS
This line(s) represent the obstruction caused by WISPR: Red was March 2013 and 2011-ish envelope and yellow is the Sept 2013 one. Outer purple is the closed door (Sept 2013).
Don’t think this red obstruction was in 2011 model. Part of WISPR?
At 0 degrees elevation, the viewing angle is 9.4 degrees off from looking straight at the sun
ion
electron

## Slide 8: EPI-Lo Wedge Components
EPI-LO INDIVIDUAL SENSOR WEDGE ASSY CROSS SECTION
BOTTOM MCP 2.9kV
TOP MCP, 900V
MCP GRID, 1kV
SCREWS, TORLON
SENSOR WEDGE COVER, GROUND
APERTURE START FOIL GRID, GROUND
INSULATOR, SENSOR WEDGE COVER
ULTEM 1000
MCP COMPRESSION PLATE, 1kV
(TOP OF ANODE PWB NOT SHOWN, 3KV)
COLLIMATORS, AT GROUND
ACCELERATION SURFACE, +1kV

## Slide 9: EPI-Lo Wedge, Exploded View
EPI-LO INDIVIDUAL SENSOR WEDGE ASSY CROSS SECTION
APERTURE START FOILS, GROUND
WEDGE COVER INSULATOR, ULTEM 2300
WEDGE GRID, 1KV
WEDGE COVER, GROUND
SIDE WALL REMOVED FOR CLARITY

## Slide 10: Sensor Functional Schematic
First
  xxxxx
  xxxxxx
Second
Third
  xxxxxx
  xxxx
Forth
  xxxxxxx
  xxxxxx
Microchannel
Plate (MCP)
Anode
Solid State
Detectors (SSDs)
Start Foils
Stop Foil
Start Foils
Collimators
Baffles
MCP Grid
Start Acceleration
Grid
Stop Accel. Grid
MCP Spacer

## Slide 11: Sensor Voltages
Ground
1000 V
Equipotential
Volume
900 V
2900 V
3000 V
start e-
acceleration
region
MCP
Anode
Ground
stop e-
accel.
region
SSDs
All Foils
at Ground
100 V retarding
potential region

## Slide 12: Secondary Electrons & Potentials
100 V retarding
potential region
Equipotential
Volume
2000 V
Across Plates
100 V
MCP-to-Anode
Grid
Grid
Grid
SSDs
start e-
acceleration
region
stop e-
accel.
region
MCP
Anode

## Slide 13: Ions (Energy, TOF, and Position)
Incident Ion
Secondary
Electrons
Start
Stop
MCP
Anode
SSDs

## Slide 14: Electrons: Energy (& Position)
Secondary
Electrons
(infrequent)
MCP
Anode
SSDs
Incident Electron

## Slide 15: Energetic Electron Measurement
Secondary Electrons from Start Foils possible, but Low Probability;
Start Electron present identifies entrance aperture
?
?
?
?
?

## Slide 16: Energetic Ion Measurement
Ion Measurement Logic can require TOF or not, but with not TOF, no species

## Slide 17: Electrostatic Optics
SIMION simulation of electron optics using ~3eV random initial electron velocities.

## Slide 18: Sensor Voltages
Ground
1000 V
Equipotential
Volume
900 V
2900 V
3000 V
start e-
acceleration
region
MCP
Anode
Ground
stop e-
accel.
region
SSDs
All Foils
at Ground
100 V retarding
potential region

## Slide 19
Elevation 1
Elevation 2
Elevation 3
Elevation 4
Elevation 5
Name | Mean TOF (ns) | TOF Sigma (ns)
Elevation 1 | 1.06 | 0.19
Elevation 2 | 2.17 | 0.13
Elevation 3 | 2.35 | 0.05
Elevation 4 | 2.33 | 0.06
Elevation 5 | 2.32 | 0.12
Stops | 5.23 | 0.11
Initial Energy Distribution
Electron Dispersion
Hit Locations
Sample Trajectories
SIMION, Alegrini Energy Distribution
19

## Slide 20: EPI-Lo Wedge, Exploded View
EPI-LO INDIVIDUAL SENSOR WEDGE ASSY CROSS SECTION
APERTURE START FOILS, GROUND
WEDGE COVER INSULATOR, ULTEM 2300
WEDGE GRID, 1KV
WEDGE COVER, GROUND
SIDE WALL REMOVED FOR CLARITY

## Slide 21: MCP Inter-Plate Spacer also a mask
EXPLODED VIEW “NEW” MCP ASSEMBLY
2.9KV LUG ASSY
900 V LUG ASSY
1KV LUG ASSY
GUARD WASHER
SCREW
WASHER
LINEAR EXPANDERS, 3 PLACES
2.9KV ELECTRODE
SLIDES FOR EXPANDERS, 2 PLACES
SLIDE-SPACER-ELECTRODE
MCP
SPACER .002
ALTERNATE SPACER
MCP
900V ELECTRODE
COMPRESSION PLATE
INSULATED SHIM, ULTEM 1000
NUT
SCREW

## Slide 22
Elevation 1
Elevation 2
Elevation 3
Elevation 4
Elevation 5
Name | Mean TOF (ns) | TOF Sigma (ns)
Elevation 1 | 1.06 | 0.19
Elevation 2 | 2.17 | 0.13
Elevation 3 | 2.35 | 0.05
Elevation 4 | 2.33 | 0.06
Elevation 5 | 2.32 | 0.12
Stops | 5.23 | 0.11
Initial Energy Distribution
Electron Dispersion
Hit Locations
Sample Trajectories
SIMION, Alegrini Energy Distribution
22

## Slide 23: Microchannel Plate (MCP)

## Slide 24: Mass Resolution

## Slide 25: Mass Resolution

## Slide 26: Collimators and Start Foils
Foils Mount on Apertures (red)
  Each elevation tailored for equal GF
  Second foil planned at intermediate baffle
Collimators screw-mount to outer cover, capture foils
  Each Elevation and azimuth is unique

## Slide 27: EPI-Lo SSD Assemblies (8)
SSD ASSY TO WEDGE EXPLODED CROSS SECTION
ENERGY PWB ASSY
SSD ASSY
SSD ASSY HOUSING, ULTEM 1000
SSD WEDGE GRID, 1KV
SSD STOP GRID-FOIL, GROUND
SSD INSULATOR, ULTEM 1000
SCREWS, TORLON
GROUNDING SCREW
SENSOR WEDGE 1KV

## Slide 28: Anti-coincidence System
Electron SSD is backed by an anti-coincidence SSD

## Slide 29: Anti-coincidence System: GEANT
Setup for GEANT4 model of electron SSD anticoincidence

## Slide 30: EPI-Lo Block Diagram

## Slide 31: Sensor Voltages
MCP
Anode/Preamp
MCP
+2900V
HVPS Out +3000V
+ 900V
MCP Return is a ground reference at HVPS (allows MCP current measurement across resistor)
SSD
900V (total diode drop)
100
+3000V
1000V
Grounded Foils
Collimators
MCP return
Tecan Grid
100V
Voltage Control Cartoon

## Slide 32: EPI-Lo Block Diagram

## Slide 33: Light & Dust Mitigation
TBD

## Slide 34: Dust: Simulating the Environment
+Z-ish
45 deg
90 deg
135 deg
180 deg
225 deg
270 deg
315 deg
azimuthal angle
0 deg
0 deg
22.5 deg
45 deg
67.5 deg
90 deg
+Z
+x-ish
elevation angle
align with plane of instrument
Virtual surface
Back foil
Collimator
opening
Impacts within the back foil angle will pass through the collimator opening and front foil and hit the back foil
Front foil
Impacts within the front foil angle will pass through the collimator opening and hit the front foil
Impacts in the hemisphere will hit the “virtual surfaces” (ignores the collimator)
NOTE: actually only polar angle is limited, not position on surface, so foil area was adjusted to account for this and produce actual foil geometry factors,

## Slide 35: Dust: Simulating the Environment
Each unprotected foil would expect ~10 damage-inducing hits per mission. Reduced to ~1 hit per mission with collimator. Order of magnitude decrease.

## Slide 36: Dust: Lab tests
TBD

## Slide 37: SLIDE PREPARATION NOTES
Are we handling dust in this presentations? Assuming yes.
Should we discuss photons suppression here? Assuming yes, it’s background.
How about details like how the MCPs work (in general)? Or this could be handled when someone discusses the anode boards. Special features like the inter-plate mask (or is that handled in a mechanical section)?
Backup slides: MCPs
We are only supposed to respond to requirements, so I wasn’t going to go into any of the rate estimates from the work that I (Matt) and Rob D. did since that is for goals. Is this the right approach?

## Slide 38: Follow Up from Peer Reviews
TBD

## Slide 39: Summary
TBD

## Slide 40: Backup

## Slide 41: Collimator Foil FOV
Foil FOVs shown in yellow (these are not with the latest collimators).

## Slide 42: EPI-Lo Instrument Peer Review
EPI-LO INDIVIDUAL SENSOR WEDGE ASSY CROSS SECTION
APERTURE START FOILS, GROUND
WEDGE COVER INSULATOR, ULTEM 2300
WEDGE GRID, 1KV
WEDGE COVER, GROUND
SIDE WALL REMOVED FOR CLARITY

## Slide 43: EPI-Lo Instrument Peer Review
EXPLODED VIEW “NEW” MCP ASSEMBLY
2.9KV LUG ASSY
900 V LUG ASSY
1KV LUG ASSY
GUARD WASHER
SCREW
WASHER
LINEAR EXPANDERS, 3 PLACES
2.9KV ELECTRODE
SLIDES FOR EXPANDERS, 2 PLACES
SLIDE-SPACER-ELECTRODE
MCP
SPACER .002
ALTERNATE SPACER
MCP
900V ELECTRODE
COMPRESSION PLATE
INSULATED SHIM, ULTEM 1000
NUT
SCREW

## Slide 44: Principles of Operation